NV Quenching Readout System

Quick Reference — Firmware & Software

Samuel Sanchez  |  Fraunhofer IAO, Quantum Open Innovations  |  2025

# At a Glance

| 16-bit ADC resolution | 7.8 µV/LSB At GAIN_SIXTEEN | 0.34 mV RMS System noise | 2–3 % Quench amplitude |
| --- | --- | --- | --- |

| ~86 Hz Step rate | ~22 Hz WebSocket rate | ~91 s CSV buffer | < 20 € Total BOM cost |
| --- | --- | --- | --- |

Figure 1 Live interface with built in settings panel and usage guide

# System Overview

| Hardware chain - BPW34 photodiode — light-to-current - MCP607 TIA — Rf = 1 MΩ, 0–1 V output - RC filter — 2-stage, fc ≈ 160 kHz - ADS1115 — 16-bit I²C ADC, 860 SPS - ESP32-S3 — M5 Stamp S3, Wi-Fi AP Connectivity - Wi-Fi AP — SSID: NV-Plotter-N - IP — 192.168.10.1 - mDNS — http://nv-plotter.local - HTTP :80 — browser app + CSV export - WebSocket :81 — live data + config | Signal path - Average N fresh ADC reads - Display IIR filter — slow (α configurable) - Detection IIR filter — fast (α = 0.50, fixed) - Baseline subtraction — delta = filtered − baseline - Hysteresis comparator — entry + 50 ms guard - Dip latch — accumulates over 4-step window - Publish JSON — ~22 Hz to all WS clients - CSV ring buffer — 2000 entries, ~91 s Firmware modules - nv_signal — ADC, filters, dip, RPM - csv_logger — buffer + HTTP export - web_ui — AP, HTTP, WebSocket - web_page — full browser app (PROGMEM) - main — loop, decimation, latch |
| --- | --- |

# System Performance

| Parameter | Value | Notes |
| --- | --- | --- |
| ADC resolution | 16-bit (7.8 µV/LSB) | GAIN_SIXTEEN, ±0.256 V FS |
| Noise single read | ~2 LSB RMS ≈ 15.6 µV | ADS1115 datasheet at 860 SPS |
| Noise avg = 10 | ~0.6 LSB RMS ≈ 4.7 µV | √10 improvement |
| System noise (measured) | 0.34 mV RMS | At TIA output |
| Baseline signal level | 0.80–1.00 V | 525-532 nm LED, BPW34, Rf = 1 MΩ |
| Quench dip amplitude | 15–30 mV (2–3%) | 1920-3840 ADC counts |
| Display filter −3 dB | ~0.69 Hz | α = 0.05, ~86 steps/s |
| Detection filter −3 dB | ~21 Hz | α = 0.50, ~86 steps/s |
| Step rate | ~86 steps/s | 860 SPS / avg = 10 |
| WebSocket rate | ~22 Hz | Step rate / SEND_EVERY_N = 4 |
| CSV buffer | 2000 samples (~91 s) | Static allocation, ~56 KB RAM |
| Max dip rate | 1200 dips/min | MIN_DIP_INTERVAL_MS = 50 ms |
| Encoder — mean RPM | 61.5 rpm (±3.9 rpm, 6.3%) | 6 dips; theoretical 58.6 rpm |
| Encoder — deviation | +4.9 % | SG90 servo dynamics, not ADC error |
| Total BOM cost | ~20 € | Including ADS1115 and Fabricated PCB |

# Configuration Parameters

## Run-time  (adjustable from browser, no reflash)

| Parameter | Default | Range | Unit | Effect |
| --- | --- | --- | --- | --- |
| IIR α — display filter | 0.05 | 0.01–0.50 | — | Displayed trace smoothness. No effect on dip detection. |
| Averaging depth | 10 | 1–32 | samples | Noise reduction (√N). Adds N × 1.16 ms latency per step. |
| Detection threshold | −15 | −5000–0 | ADC cts | Dip trigger level. 1 ct ≈ 7.8 µV. Set to ~50% of dip depth. |

## Compile-time  (edit source, reflash)

| Constant | File | Default | Description |
| --- | --- | --- | --- |
| SEND_EVERY_N | main.cpp | 4 | Decimation ratio. 1 of N steps forwarded to WS and CSV. |
| RPM_HIST | nv_signal.cpp | 8 | Inter-dip interval ring buffer depth. Larger = smoother RPM. |
| MIN_DIP_INTERVAL_MS | nv_signal.cpp | 50 ms | Dead time between dip entries. 50 ms ↔ max 1200 dips/min. |
| DETECT_ALPHA | nv_signal.cpp | 0.50 | Detection filter α. Fixed — do not reduce below 0.30. |
| CSV_BUF_N | csv_logger.cpp | 2000 | Ring buffer capacity. Each entry ≈ 28 bytes. |

Threshold calibration 1. Set Mode to Δ. 2. Approach the magnet and read the minimum delta in the metric strip. 3. Set threshold to ~50% of that minimum. Example: delta reaches −80 → set threshold to −40. 4. Press Apply Settings. Confirm red dip markers appear on approach.

# Main Features

| Signal Processing - Dual IIR filters — independent display and detection paths - Paced averaging — N distinct conversions per step - Static baseline — 60-sample average at startup - Live recalibration — /recal HTTP or browser button - Hysteresis comparator — 50% exit threshold - Dip guard timer — prevents burst false triggers - RPM averaging — 8-interval ring buffer - Angular velocity & accel. — ω and α computed per dip Data Export - CSV download — /export.csv, chunked HTTP - Relative timestamps — seconds from oldest sample - Columns — t_s, raw, filtered, delta, dip, rpm - Auto filename — nv_export.csv via Content-Disposition | Browser Interface - Zero dependencies — served from ESP32 flash (PROGMEM) - Live metric strip — raw, filt, Δ, RPM, dip, mode - 5000-point ring buffer — typed arrays, no GC pressure - 60 fps canvas render — signal trace + dip markers - HiDPI aware — crisp on Retina / AMOLED - Mobile responsive — adaptive layout below 420 px - Freeze / Clear — view control without data loss - Autoscale Y — or manual zoom via pinch / Shift+scroll Interaction Gestures - 1-finger drag — pan time axis - 2-finger pinch (H) — zoom time axis - 2-finger pinch (V) — zoom Y axis - Mouse drag — pan - Scroll wheel — zoom time axis - Shift + scroll — zoom Y axis |
| --- | --- |

# HTTP & WebSocket Quick Reference

| Endpoint | Method | Description |
| --- | --- | --- |
| 192.168.10.1 or nv-plotter.local | GET | Browser application (HTML/CSS/JS, ~12 KB, no CDN) |
| /export.csv | GET | Download full ring buffer as CSV (chunked, nv_export.csv) |
| /recal | GET | Re-run 60-sample baseline capture; broadcasts recal_done to all WS clients |
| ws://192.168.10.1:81/ | WebSocket | Live data stream at ~22 Hz; accepts cfg messages; replies cfg_ack |

## WebSocket messages

| Server → Client (sample, ~22 Hz) { "t": 45231, // ms from boot "raw": 18432.5, // avg ADC counts "filtered": 18431.1, // display IIR "delta": -42.3, // counts vs baseline "dip": 1, // 1 on entry only "rpm": 61.5 } | Client → Server (config update) { "type": "cfg", "alpha": 0.05, "avg": 10, "thr": -45.0 } Server → Client (ack / recal) { "type": "cfg_ack", "alpha": 0.050, "avg": 10, "thr": -45.00 } { "type": "recal_done" } |
| --- | --- |

# Build & Connect

| platformio.ini [env:m5stack-stamps3] platform = espressif32@>=6 board = m5stack-stamps3 framework = arduino lib_deps = adafruit/Adafruit ADS1X15@^2.6.0 Flash command pio run --target upload | First-use steps - Connect to Wi-Fi: NV-Plotter-2 / 12345678 - Open: http://nv-plotter.local - Observe live delta in the metric strip - Approach magnet → note minimum delta - Set threshold to ~50% of that value → Apply - Confirm red dip markers on approach - Recalibrate if delta drifts from zero |
| --- | --- |