Fraunhofer IAO

Firmware & Software Documentation

Suplement to: NV_Samuel_Report

Samuel Sanchez 2-19-2026

# Introduction

This document provides the technical reference for the embedded firmware and browser-based visualization software of the NV-diamond fluorescence quenching readout prototype. It describes the final implemented system: the signal-processing architecture, module responsibilities, configurable parameters, data structures, communication protocol, and measured performance.

The system runs on an M5 Stamp S3 (ESP32-S3) microcontroller. It acquires raw photodiode counts from a transimpedance amplifier (TIA) connected to a BPW34 photodiode, applies real-time digital filtering and baseline subtraction, detects fluorescence quenching dip events caused by an external magnetic field, estimates rotational parameters (RPM, angular velocity, angular acceleration), streams live data to a browser over Wi-Fi, and provides timestamped CSV export for offline analysis.

# System Architecture

## Module Overview

The firmware is structured as five independent modules with clearly separate responsibilities. This structure allows signal-processing parameters to be tuned independently of the communication layer.

| File | Responsibility | Key Dependencies |
| --- | --- | --- |
| nv_signal.cpp / .h | ADC acquisition, IIR filtering, baseline subtraction, dip detection, RPM estimation | Adafruit ADS1X15, Wire (I²C) |
| csv_logger.cpp / .h | Circular sample buffer; chunked HTTP CSV export with relative-second timestamps | WebServer |
| web_ui.cpp / .h | Wi-Fi Access Point, HTTP server (port 80), WebSocket server (port 81), configuration handler | WiFi, WebServer, WebSocketsServer, ESPmDNS |
| web_page.cpp / .h | Complete single-page browser application stored in flash (PROGMEM); served as a static string | None (self-contained) |
| main.cpp | Main acquisition loop; dip event accumulation across the decimation window; sample forwarding to logger and WebSocket | All modules |

## Signal Path — From Photon to Browser

The table below traces the complete measurement chain from the photodiode to the rendered waveform in the browser, identifying the location and nature of each transformation.

| Stage | Location | Operation | Output |
| --- | --- | --- | --- |
| Photon → Current | BPW34 photodiode | Light-to-current conversion | i⁐⁄ (nA–µA) |
| Current → Voltage | MCP607 TIA (Rf = 1 MΩ) | Transimpedance gain | Vₜᴵᴬ (0–1 V typ.) |
| Anti-alias filtering | RC network (100 Ω, 10 nF ×2) | 2-stage passive low-pass, f⁣ ≈ 160 kHz | Filtered analogue voltage |
| Digitisation | ADS1115, GAIN_SIXTEEN, 860 SPS | 16-bit ADC; 1 LSB = 7.8125 µV at ±0.256 V FS | Raw counts (int16_t) |
| Hardware averaging | nv_step() — readOneFreshConversion() | N distinct conversions summed and divided; paced at the ADC conversion period | rawAvg (float counts) |
| Display IIR filter | nv_step(), configurable α | filteredRaw += α·(rawAvg − filteredRaw) | filteredRaw (float) |
| Detection IIR filter | nv_step(), fixed α = 0.50 | detectedFiltered += 0.50·(rawAvg − detectedFiltered) | detectedFiltered (float) |
| Baseline subtraction | nv_step() | delta = filteredRaw − baseline | delta (float counts) |
| Dip detection | nv_step() — hysteresis comparator | Entry if detectionDelta < threshold and inter-dip guard satisfied | dipEvent (bool), RPM (float) |
| Decimation & latch | loop() in main.cpp | Accumulate dipEvent flag over 4 steps; forward one sample per 4 steps | Latched NvSample |
| WebSocket publish | web_publish_sample() | JSON frame broadcast to all clients at ~22 Hz | JSON string (~200 B) |
| CSV logging | csv_push() | Write to 2000-entry circular ring buffer | Ring buffer entry |
| Browser rendering | JavaScript canvas loop | 5000-point ring buffer; 60 fps canvas draw; pan and pinch-zoom | Live waveform display |

# Module Reference

## nv_signal — Signal Acquisition and Processing

This module is the scientific core of the firmware. Every decision that affects measurement fidelity — ADC gain, averaging depth, filter bandwidth, detection sensitivity, and RPM accuracy — is implemented and encapsulated here. The public API consists of five functions: nv_init(), nv_recalibrate(), nv_step(), nv_get_config(), and nv_set_config().

### ADS1115 Configuration

Gain: GAIN_SIXTEEN — full-scale range ±0.256 V, 1 LSB = 7.8125 µV. Chosen to maximise resolution for the small TIA output variations (15–30 mV) that correspond to 2–3% fluorescence quenching. The gain setting keeps the expected baseline voltage of ~0.5 V well within the input range.

Data rate: RATE_ADS1115_860SPS — the highest available rate, giving one conversion every ≈1163 µs. The faster rate is compensated by the averaging stage.

Conversion mode: Continuous (startADCReading, continuous = true). The ADC free-runs and always presents the most recent result in its output register, which is read non-blocking via getLastConversionResults().

### Baseline Capture

The baseline is captured once at startup and can be refreshed at any time by calling nv_recalibrate() — available via the browser Recalibrate button or the /recal HTTP endpoint — without rebooting the device. The capture procedure:

- Discard 10 warm-up conversions at the paced rate to allow the ADC reference and TIA output to settle.

- Accumulate 60 consecutive fresh conversions (~70 ms total); compute their arithmetic mean as the static baseline.

- Initialise both IIR filter states (filteredRaw and detectedFiltered) to the baseline value to prevent transient false events on the first acquisition steps.

- Reset all dip-detection state: dipState = false, lastDipTimeMs = 0, RPM history cleared to zero.

Baseline stability The baseline is a static scalar. Slow thermal drift of the TIA or gradual LED power changes will cause the delta signal to walk from zero over time. If this is observed, pressing Recalibrate in the browser re-runs the 70 ms capture procedure and updates all filter states without interrupting the Wi-Fi connection.

### Hardware Averaging

nv_step() averages samplesToAvg distinct ADC conversions per processing step. Each read is paced to the ADS1115 conversion period so that every call to readOneFreshConversion() retrieves a genuinely new result. The pacing function compares the elapsed microseconds since the last read against CONV_PERIOD_US (1163 µs at 860 SPS) and inserts a brief wait if needed.

nv_signal.cpp  —  readOneFreshConversion()

static inline int16_t readOneFreshConversion() {

static unsigned long lastReadUs = 0;

unsigned long now = micros();

if (now - lastReadUs < CONV_PERIOD_US)

delayMicroseconds(CONV_PERIOD_US - (now - lastReadUs));

lastReadUs = micros();

return ads.getLastConversionResults();}

With samplesToAvg = 10 and GAIN_SIXTEEN, the theoretical noise reduction from averaging is √10 ≈ 3.16×, reducing effective input-referred noise from ~2 LSB RMS (ADS1115 datasheet) to ~0.6 LSB RMS ≈ 4.7 µV.

### Dual-Filter Architecture

Two independent first-order IIR low-pass filters operate on the same average raw signal. Their roles are strictly separated to decouple visual smoothness from detection responsiveness.

| Filter | State variable | α | Time constant (τ) | Role |
| --- | --- | --- | --- | --- |
| Display filter | filteredRaw | Configurable 0.01–0.50 | ≈1/α steps | Produces the smooth waveform rendered in the browser. Default α = 0.05 gives τ ≈ 20 steps (≈0.23 s at 86 steps/s). Not used for threshold decisions. |
| Detection filter | detectedFiltered | Fixed 0.50 | ≈2 steps | Tracks sharp dip edges fast enough to follow real quench events. Feeds the threshold comparator and hysteresis exit test exclusively. |

The IIR update equation for both filters is identical in form:

y[n] = y[n−1] + α · (x[n] − y[n−1])

where x[n] is the current averaged raw count and y[n] is the filter output. For the display filter with α = 0.05 and a step rate of ~86 Hz, the −3 dB cut-off frequency is approximately f⁣ = α·f_step / (2π) ≈ 0.69 Hz.

### Dip Detection

Dip detection uses a two-threshold (hysteresis) state machine operating on the fast detection filter output. The detection delta is defined as detectionDelta = detectedFiltered − baseline.

| Transition | Condition | Guard |
| --- | --- | --- |
| Armed → Dip | detectionDelta < DIP_THRESHOLD | Minimum inter-dip interval of 50 ms must have elapsed since the previous confirmed dip (prevents burst re-triggering near the exit hysteresis). |
| Dip → Armed | detectionDelta > DIP_THRESHOLD × 0.5 | None — exit is immediate once the fast filter recovers above the hysteresis level. |

On each Dip entry: dipEvent is set to true for that single step, and computeRpm() is called with the current millisecond timestamp. The 50 ms minimum inter-dip guard (MIN_DIP_INTERVAL_MS) corresponds to a maximum detectable dip rate of 1200 events per minute, well above any rotation speed tested with the prototype.

nv_signal.cpp  —  threshold comparator in nv_step()

if (!dipState

&& detectionDelta < DIP_THRESHOLD

&& (lastDipTimeMs == 0 ||

(now - lastDipTimeMs) >= MIN_DIP_INTERVAL_MS)) {

dipState = true;  dipEvent = true;

computeRpm(now);}

if (dipState && detectionDelta > DIP_THRESHOLD * 0.5f)

dipState = false;}

### RPM and Angular Velocity

Each confirmed dip entry records its millisecond timestamp. The inter-dip interval dt is stored in a circular ring buffer of depth RPM_HIST = 8. The mean of all stored intervals is used to compute RPM, smoothing out single-period timing jitter:

RPM = 60 000 / mean(dt₁, dt₂, …, dt₈)

Angular velocity is derived as ω = RPM × 2π / 60 (rad/s). Angular acceleration α = Δω / Δt (rad/s²) is computed between consecutive dip events and reported in the serial encoder output. These are the values listed in the Results section of the main report.

## csv_logger — Sample Buffer and CSV Export

csv_logger maintains a static circular buffer of 2000 NvSample entries allocated in BSS (zero-initialised, no heap allocation). Samples are written by csv_push() on every forwarded step. When the buffer is full, new entries overwrite the oldest.

| Buffer parameter | Value | Notes |
| --- | --- | --- |
| Capacity | 2000 samples | Static allocation — no heap fragmentation risk |
| Entry size | ~28 bytes | uint32_t + float×4 + bool + float |
| Total RAM used | ~56 KB | Well within the ESP32-S3's 520 KB SRAM |
| Write rate | ~22 Hz | 860 SPS / avg=10 / SEND_EVERY_N=4 |
| Buffer duration | ~91 s | 2000 samples / 21.5 Hz |

CSV export is delivered as a chunked HTTP response (CONTENT_LENGTH_UNKNOWN) so that the entire payload is never buffered in heap memory simultaneously. Rows are formatted and streamed one at a time with snprintf(). Timestamps are expressed as seconds elapsed from the oldest sample in the export window, making the data immediately plottable in any spreadsheet application.

csv_logger.cpp  —  relative timestamp conversion

uint32_t t0 = buf[start % CSV_BUF_N].t_ms;     // oldest entry

float t_s = (float)(s.t_ms - t0) / 1000.0f;    // seconds from start

The CSV column order is: t_s, raw, filtered, delta, dip (0 or 1), rpm. A Content-Disposition header is included so that browsers automatically suggest the filename nv_export.csv on download.

## web_ui — Wi-Fi, HTTP, and WebSocket Server

The ESP32-S3 operates as a Wi-Fi Access Point (SSID: NV-Plotter-N, password: 12345678) with a fixed IP address of 192.168.10.1. This removes any dependency on laboratory network infrastructure and allows direct connection from a smartphone or laptop without router access. mDNS is advertised under the hostname nv-plotter, enabling connection via http://nv-plotter.local.

| Endpoint | Protocol | Description |
| --- | --- | --- |
| GET / | HTTP port 80 | Serves the complete browser application from PROGMEM. Cache-Control: no-store prevents stale caching across firmware updates. |
| GET /export.csv | HTTP port 80 | Triggers chunked CSV streaming of the 2000-entry ring buffer. Content-Disposition header set for automatic browser download. |
| GET /recal | HTTP port 80 | Executes nv_recalibrate() (~70 ms), then broadcasts {"type":"recal_done"} to all connected WebSocket clients. Returns 200 OK. |
| WS port 81 | WebSocket | Bidirectional channel. Server pushes JSON sample frames at ~22 Hz. Client sends configuration updates; server replies with clamped values. See Appendix B for the full message protocol. |

Configuration messages received via WebSocket are parsed by extractNumber(), a minimal hand-written function that locates a named key and reads the following numeric token. This avoids the overhead of a general-purpose JSON library (ArduinoJson) while remaining sufficient for the fixed, schema-controlled message format.

## web_page — Browser Application

The complete user interface is stored as a C string constant in the ESP32's flash (PROGMEM) and served on demand. It requires no internet connection, no CDN, and no external dependencies. The HTML, CSS, and JavaScript fit within approximately 12 KB.

### Metric Strip and Settings Panel

On load, a six-cell metric strip at the top of the page shows live values for raw ADC count, filtered ADC count, delta (baseline-subtracted), RPM, dip indicator, and the active display mode selector. Below it, a full-width settings panel is open by default, exposing all three configurable signal-processing parameters as range sliders with companion labels and a numeric text input for precise threshold entry.

### Client-Side Data Buffer

The browser maintains a ring buffer of 5000 points using three typed arrays: Float64Array for timestamps, Float32Array for signal values, and Uint8Array for dip flags. Typed arrays are used deliberately — they are contiguous in memory and avoid garbage-collector pressure during the 60 fps draw loop. The ring buffer correctly handles the wrap-around index mapping in both the partially filled and fully overwritten states.

### Canvas Rendering

- HiDPI support: the canvas pixel buffer is allocated at cssW × devicePixelRatio. ctx.setTransform(dpr, 0, 0, dpr, 0, 0) scales all draw calls to CSS-pixel coordinates, producing a sharp trace on Retina and high-DPI AMOLED displays without changing any rendering logic.

- Adaptive layout: on screens narrower than 420 px, Y-axis margin narrows from 62 px to 48 px, Y-tick count reduces from 6 to 4, and X-tick count from 5 to 3 — maintaining readability on 375 px portrait phone screens.

- Signal trace: drawn as a continuous polyline in cyan (#5fd7ff) at 2 px width (1.5 px on narrow screens).

- Dip markers: each sample with dip = 1 produces a full-height vertical red line at the corresponding X position. Markers are rendered in a second pass after the signal trace to guarantee visibility.

- Frozen mode: when active, the view stops following new data. A translucent amber overlay and FROZEN label are drawn over the plot. New data continues to accumulate in the ring buffer so no samples are lost.

- Autoscale Y: when checked, the Y axis range recomputes on every frame from the visible window's min/max, with 15% padding. When unchecked, the manual yCenter and ySpan values are used, adjustable by Shift+scroll or vertical pinch.

### Interaction Gestures

| Input | Device | Action |
| --- | --- | --- |
| 1-finger drag | Touch | Pan the time axis — viewEnd shifts by dx × (viewSpan / cssW) logical points |
| 2-finger horizontal pinch | Touch | Zoom the time axis, anchored at the midpoint between the two contact points |
| 2-finger vertical pinch | Touch | Zoom the Y axis (only when Autoscale Y is disabled) |
| Click and drag | Mouse | Pan — identical formula to 1-finger drag |
| Scroll wheel | Mouse | Zoom the time axis, anchored at the cursor X position |
| Shift + scroll wheel | Mouse | Zoom the Y axis (only when Autoscale Y is disabled) |

Touch events are tracked by pointer identifier in an activeTouches dictionary. Horizontal versus vertical pinch is determined by comparing |dx| and |dy| between the two contact points: horizontal dominance routes to time-axis zoom; vertical dominance routes to Y-axis zoom.

### User Instructions Panel

A permanent instructions section below the plot describes all gestures, button functions, and slider effects in plain language accessible to users ranging from high-school students to research engineers. It includes a dedicated threshold calibration procedure and explains the unit convention (ADC counts, 1 count ≈ 7.8 µV at GAIN_SIXTEEN).

## main — Acquisition Loop and Dip Accumulation

The main loop calls nv_step() on every iteration. One out of every SEND_EVERY_N = 4 steps is forwarded to the WebSocket publisher and CSV logger, reducing network traffic to ~22 Hz.

Because NvSample.dip is a single-shot pulse — true for exactly the one step where the detection threshold is first crossed — a static accumulation latch (dipPending) ensures that a dip event occurring on any of the 4 steps in a decimation window is reliably carried through to the forwarded sample.

main.cpp  —  acquisition loop with dip accumulation latch

static bool dipPending = false;

void loop() {

web_loop();

NvSample s = nv_step();

if (s.dip) dipPending = true;

if (++counter % SEND_EVERY_N == 0) {

s.dip     = dipPending;

dipPending = false;

csv_push(s);

web_publish_sample(s.raw, s.filtered, s.delta, s.dip, s.rpm);

}

}

# Configuration Parameters

Three parameters can be adjusted at run-time from the browser settings panel without reflashing. Changes are applied immediately on pressing Apply Settings; the WebSocket cfg_ack response confirms the values in use after internal clamping.

| Parameter | Default | Range | Unit | Effect |
| --- | --- | --- | --- | --- |
| IIR α (display filter smoothing) | 0.05 | 0.01–0.50 | — | Bandwidth of the displayed trace. Lower α = smoother, slower-responding curve. Higher α = closer to raw signal, more visible noise. Does not affect dip detection. |
| Averaging depth | 10 | 1–32 | samples | Number of distinct ADC conversions averaged per processing step. Improves input-referred noise by √N. Also increases step latency by N × 1163 µs. |
| Detection threshold | −15 | −5000 to 0 | ADC counts | Level that detectionDelta must fall below to register a dip. 1 count = 7.8125 µV at GAIN_SIXTEEN. Adjustable via slider (step 5) or direct numeric entry. |

Threshold calibration 1. Set the Mode selector to Δ (delta). 2. Perform a magnet approach while observing the delta value in the metric strip. 3. Note the minimum (most negative) delta reached during the dip. 4. Set the threshold to approximately 50–60 % of that minimum. Example: if delta reaches −80 counts, set threshold to −40 to −48 and press Apply Settings. 5. Confirm that red dip markers appear on subsequent approaches and disappear on removal.

Three additional parameters are compile-time constants in the source files:

| Constant | File | Default | Description |
| --- | --- | --- | --- |
| SEND_EVERY_N | main.cpp | 4 | Decimation ratio. Every Nth step is forwarded to WebSocket and CSV. Adjust to trade off temporal resolution against network load. |
| RPM_HIST | nv_signal.cpp | 8 | Depth of the inter-dip interval ring buffer. Larger values give smoother RPM at the cost of slower response to genuine speed changes. |
| MIN_DIP_INTERVAL_MS | nv_signal.cpp | 50 ms | Minimum time between consecutive dip entries. Prevents burst false-triggers near the hysteresis exit level. 50 ms ↔ max 1200 dips/min. |

# Data Structures

## NvSample

NvSample is the fundamental unit of data passed between all modules. It is populated by nv_step(), written to the CSV ring buffer, and serialised to JSON for WebSocket transmission.

nv_signal.h  —  NvSample

struct NvSample {

uint32_t t_ms;       // millis() timestamp at end of step

float    raw;        // averaged raw ADC count (pre-filter)

float    filtered;   // display IIR output (slow, α configurable)

float    delta;      // filtered − baseline  (ADC counts)

bool     dip;        // true on dip ENTRY only (single-shot pulse)

float    rpm;        // moving-average RPM from last 8 inter-dip intervals

};

## NvConfig

NvConfig carries the three run-time configurable parameters between the WebSocket handler and the signal processing module.

nv_signal.h  —  NvConfig

struct NvConfig {

float    alpha;    // display IIR smoothing factor

uint16_t avg;      // number of fresh ADC conversions to average

float    thr;      // dip detection threshold (negative ADC counts)

};

# System Performance

The table below consolidates measured and theoretical performance figures of the complete acquisition and detection chain as validated in the final prototype evaluation.

| Parameter | Value | Condition |
| --- | --- | --- |
| ADC resolution | 16-bit (7.8125 µV/LSB) | ADS1115, GAIN_SIXTEEN (±0.256 V FS) |
| Single-read noise (typ.) | ~2 LSB RMS ≈ 15.6 µV | ADS1115 datasheet at 860 SPS |
| Noise with avg = 10 | ~0.6 LSB RMS ≈ 4.7 µV | √10 improvement from averaging |
| Measured system noise | 0.34 mV RMS | At TIA output; dominated by optical and TIA noise, not ADC |
| Baseline signal level | 0.80–1.00 V | Typical with 525–532 nm LED, BPW34, Rf = 1 MΩ |
| Quenching dip amplitude | 15–30 mV (2–3%) | 1920–3840 ADC counts at GAIN_SIXTEEN |
| Display filter – 3 dB | ≈0.69 Hz | α = 0.05, step rate ~86 Hz |
| Detection filter – 3 dB | ≈21 Hz | α = 0.50, step rate ~86 Hz |
| Step rate | ~86 steps/s | 860 SPS / avg = 10 |
| WebSocket publish rate | ~22 Hz | Step rate / SEND_EVERY_N = 4 |
| CSV buffer duration | ~91 s | 2000 entries at 22 Hz |
| Max detectable dip rate | 1200 dips/min | Limited by MIN_DIP_INTERVAL_MS = 50 ms |
| Encoder accuracy (mean RPM) | 61.5 rpm (±3.9 rpm, ~6.3%) | 6 consecutive dips; theoretical 58.6 rpm |
| Deviation from theoretical | +4.9% | Dominated by SG90 servo mechanical dynamics |
| Total system cost | ~9–16 € | Full BOM including ADS1115 external ADC |

Encoder accuracy context The 6.3 % standard deviation in RPM measurements is consistent with the known mechanical behaviour of the SG90 servo: the programmed step/delay pattern (15° per step, 1 ms delay, 500 ms endpoint dwell) produces a theoretical period of 1.024 s (58.6 rpm), but the servo’s internal dynamics cause actual motion to deviate from the commanded profile. The NV-based encoder faithfully captures the real mechanical motion rather than the ideal script.

# Build Environment and Deployment

## PlatformIO Configuration

platformio.ini

[env:m5stack-stamps3]

platform          = espressif32@>=6

board             = m5stack-stamps3

framework         = arduino

upload_speed      = 921600

upload_protocol   = esptool

build_flags       = -DARDUINO_USB_CDC_ON_BOOT=1  -std=gnu++17  -Wall

monitor_speed     = 115200

lib_deps          = adafruit/Adafruit ADS1X15@^2.6.0

## Library Dependencies

| Library | Source | Purpose |
| --- | --- | --- |
| Adafruit ADS1X15 v2.6.0+ | PlatformIO registry | ADS1115 driver: gain, data rate, continuous mode, getLastConversionResults() |
| WebServer | ESP-IDF built-in | HTTP server with chunked transfer and custom header support |
| WebSocketsServer | Links2004/arduinoWebSockets (auto-resolved by PIO) | WebSocket server on port 81 with broadcast support |
| ESPmDNS | ESP-IDF built-in | mDNS hostname advertisement (nv-plotter.local) |
| Wire | ESP-IDF built-in | I²C master interface for ADS1115 |

## First-Use Procedure

- Connect the M5 Stamp S3 via USB-C and flash with: pio run --target upload

- Connect a Wi-Fi client (phone or laptop) to the access point NV-Plotter-N, password 12345678

- Open a browser and navigate to http://192.168.10.1 or http://nv-plotter.local.

- The Settings panel opens automatically. Observe the live delta value in the metric strip.

- Perform a magnet approach and note the minimum delta count reached during the dip.

- Set the threshold to ~50 % of that minimum value and press Apply Settings.

- Confirm that red dip markers appear during subsequent magnet approaches. RPM populates after two consecutive confirmed dips.

- Use the Recalibrate button if the delta signal drifts far from zero in the absence of a magnet.

# Limitations and Recommended Future Work

## Static Baseline

The baseline is a single scalar captured at startup or on explicit recalibration. Slow thermal drift of the TIA or LED power decay will cause the delta signal to shift over time. A tracking baseline using a second, very slow IIR (α ≈ 0.0001) would compensate automatically, with care taken to prevent it from tracking the dip signal itself during magnet approach.

## First-Order Display Filter

The display IIR is a single-pole filter. For applications requiring a sharper transition between passband and stopband, a second-order Biquad IIR or a short FIR of approximately 20 taps could be implemented within the ESP32-S3 compute budget without affecting acquisition rate.

## Blocking Acquisition Loop

The readOneFreshConversion() pacing introduces a blocking delay of N × 1163 µs per step (~11.6 ms for avg = 10). During this period web_loop() is not called, which can delay WebSocket heartbeats at high averaging depths. Restructuring acquisition as a non-blocking state machine that yields to web_loop() between reads would eliminate this constraint.

## Access Point Mode Only

The sensor operates exclusively as a Wi-Fi Access Point. Adding optional Station (STA) mode would allow the device to join an existing network and forward data to remote servers or laboratory data-collection systems without requiring a direct client connection.

## No Microwave Excitation

The firmware supports all-optical quenching sensing only. Full ODMR (Optically Detected Magnetic Resonance) requires integration of a microwave signal source, frequency sweep control, and lock-in or synchronous detection. The existing IIR filtering and dip-detection architecture is compatible with this extension; only the hardware excitation layer and synchronisation logic would need to be added.

## ADS1015 for Higher Acquisition Rate

The ADS1015 (12-bit) offers conversion rates up to 3300 SPS versus 860 SPS for the ADS1115. For applications requiring finer temporal resolution of fast rotation, the ADS1015 is a pin-compatible replacement. The firmware would require CONV_PERIOD_US to be updated to 304 µs and the gain selection revisited for the reduced dynamic range.

# Appendix A — Source File Layout

PlatformIO project structure

ESP32S3_ReadOut_Web_Server/

├── platformio.ini

└── src/

├── main.cpp          — acquisition loop, decimation, dip latch

├── nv_signal.h       — NvSample, NvConfig, public API

├── nv_signal.cpp     — ADC, dual IIR filters, baseline, dip detection, RPM

├── csv_logger.h      — csv_push(), csv_write_to_client() declarations

├── csv_logger.cpp    — 2000-entry circular buffer, HTTP CSV streaming

├── web_ui.h          — web_init(), web_loop(), web_publish_sample()

├── web_ui.cpp        — Wi-Fi AP, HTTP routes, WebSocket server

├── web_page.h        — INDEX_HTML PROGMEM declaration

└── web_page.cpp      — complete browser application as C string literal

# Appendix B — WebSocket JSON Protocol

## Server → Client — Sample Frame

Broadcast to all connected clients at ~22 Hz. Field t is milliseconds from ESP32 boot; delta is in raw ADC counts; dip is 1 on the frame where a quench event is confirmed, 0 otherwise.

Sample frame (broadcast at ~22 Hz)

{  "t":        45231,

"raw":      18432.5,

"filtered": 18431.1,

"delta":    -42.3,

"dip":      1,

"rpm":      61.5

}

## Client → Server — Configuration Update

Sent by the browser on pressing Apply Settings. All three fields must be present.

Configuration message

{  "type":  "cfg",

"alpha": 0.05,

"avg":   10,

"thr":   -45.0

}

## Server → Client — Configuration Acknowledgement

Sent only to the requesting client. Values reflect the result of nv_set_config() clamping — the browser updates its sliders to match.

Acknowledgement (unicast to requesting client)

{  "type":  "cfg_ack",

"alpha": 0.050,

"avg":   10,

"thr":   -45.00

}

## Server → All Clients — Recalibration Notification

Broadcast by the server after a successful /recal HTTP request or a WebSocket recal command. All connected browsers flash the Recalibrate button green on receipt.

Recalibration notification (broadcast)

{  "type": "recal_done"   }